## Slide 1: Design of a highly Integrated Dew Point Mirror Hygrometer
Master Thesis
Final Presentation

## Slide 2: OVERVIEW - humidity
Amount of water in a gas
Described as absolute [kg/m³] or relative [%] value (to equilibrium)
Aim of this work: Development of a Dew Point Hygrometer
  cost-efficient
  yet highly accurate
[1]
Notes: Steigen wir direkt ein. Ein Taupunkthygrometer ist ein gerät zur bestimmung der Luftfeuchtigkeit. Die Luftfeuchtigkeit beschreibt erst einmal ganz generell die Menge an Wasserdampf in Luft
Luftfeuchtigkeit: Menge an Wassermolekülen in gasförmigem Zustand in einem Gas bzw. hier in der Luft
Beschrieben als absolutes Maß bspw. In Masse pro Volumen oder relativer Wert im Verhältnis zur maximalen Luftfeuchte
Maximale Luftfeuchte ist bestimmt durch den Equilibriumdampfdruck, also dem Dampfdruck, bei welchem es zwangsläufig zur Kondensation kommt. Dieser ist hauptsächlich von der Temperatur abhängig, wird aber auch durch Druck und Partikel wie Salze beeinflusst. Der Euqilibriumdampfruck steigt mit steigender Temperatur bzw. fällt bei sinkender
Ziel der Arbeit is es einen Taupunkthygrometer zu entwickeln, welcher durch hohe Integration kostengünstig und gleichzeitig mit einer hohen Genauigkeit die Luftfeuchte bestimmen kann.

## Slide 3: OVERVIEW - HYGROMETER
Device for measurement of humidity
Different kinds: mechanical, electrical, optical, chemical
Most commonly used today:
Capacitive Sensor | Chilled Dewpoint Mirror
Capacitive measurement using a hygroscopic dielectric | Dew point temperature measurement using a cooled down mirror
+ cheap and small | + ±0.1% RH accuracy
- ±1% RH accuracy | - expensive and large
- long-term drift | - influence on measured gas
Notes: Um das Ganze mal einzuordnen, erstmal einen kurzen Überblikc über die Methoden der Bestimmung der Luftfeuchte
Zur Einordnung: Ein Hygrometer ist ein Sensor zur Bestimmung der Luftfeuchte

## Slide 4: Dew Point Mirror
Mirror integrated thermal sensors (RTD)
Peltier element for cooling
Optical readout of condensate (LED/PD)
Gas flow control
LED
Regulator
Heat Pump
Regulation
Dew Point
Readout
Gas Inlet

## Slide 5: Hardware – Mirror FPC

## Slide 6: Hardware – Mirror FPC

## Slide 7: Hardware – Mirror FPC

## Slide 8: Hardware – Mirror FPC

## Slide 9: Hardware – Mirror FPC

## Slide 10: Hardware – BAse PCB

## Slide 11: Hardware – BAse PCB

## Slide 12: Hardware – BAse PCB

## Slide 13: Hardware – BAse PCB

## Slide 14: Software
Set Initial
Conditions
Wait & Fetch
Data
Data
Transmission
Dew Point Analysis
Initial Conditions:
Determine initial readings
Set thresholds for dew presence & abscense
Sensor Readout:
I²C commands for integrated sensors
Timer on interrupt for frequency measurement of capacitor
Initiate Sensor
Readout
State & PWM
Update
State & PWM Update:
State depends on thresholds
PWM for heater and cooler depends on temperature error (PID)
Dew Point Analysis:
Smoothing and interpolation using splines
1st and 2nd order derivatives
Search for significant points
Notes: Da keine onboard analyse: thresholds empirisch so bestimmt, dass definitiv betauung bei abkühlung bzw. kondensation bei erwärmung erfolgt ist

## Slide 15: Test Setup
Temperature chamber
Humidity control by deliquescence:
Cooling cycles per measurement [°C/s]: -1.3, -0.4, -0.2, -0.1
Sampling frequency: 5 Hz
temperature chamber
sealed box
sensor
fans
solution
microcontroller
data &
power
T [°C] | MgCl2 | Mg(NO3)2 * 6H20 | NaCl | K2SO4
10 | 33.47 | 57.36 | 75.67 | 98.18
25 | 32.78 | 52.89 | 75.29 | 97.59
40 | 31.60 | 48.42 | 74.68 | 96.71
[2]

## Slide 16: Single Measurement
T=25°C / RH=52.89%

## Slide 17: Single Measurement
T=25°C / RH=52.89%

## Slide 18: Single Measurement
T=25°C / RH=52.89%
ddY/dt

## Slide 19: Single Measurement
T=25°C / RH=52.89%
dY/dt

## Slide 20: Single Measurement
T=25°C / RH=52.89%

## Slide 21: Single Measurement
T=25°C / RH=52.89%

## Slide 22: All MeasurementS – Comparison

## Slide 23: T=40°C / RH=48.4%

## Slide 24: RESULTS

## Slide 25: Conclusion
Platform for dew point analysis
Simple mechanical and electronic design
Accuracy:
  Similar between IDE capacitor and proximity sensors
  Competitive with common integrated capacitive humidity sensors
  Order of magnitude less accurate than commercial dew point mirror sensors
Outlook:
  Realtime dew point analysis for feedback loop
  Data-driven offset reduction
  Test setup (constant humidity/temperature)

## Slide 26: SOURCES
[1] http://www.atmo.arizona.edu/students/courselinks/fall15/atmo170a1s2/lecture_notes/humidity/condensation_nuclei.html
[2]
L. Greenspan: Humidity Fixed Points of Binary Saturated Aqueous Solutions. Journal of Research of the National Bureau of Standards – A. Physics and Chemistry, Vol. 81A, No. 1, January-February 1977

## Slide 27: Discussion
Possible error sources:
  Realtime dew point analysis for feedback loop
  Data-driven offset reduction
  Test setup (constant humidity/temperature)
  Temperature dependencies in relaxation oscillator
  Uneven temperature distribution

## Slide 28: Discussion
Possible error sources:
  Missing feedback loop
  Missing data-driven offset reduction
  Test setup (constant humidity/temperature)
  Nonlinearities in proximity sensors
  Temperature dependencies in relaxation oscillator
  Uneven temperature distribution

## Slide 29: Software - Analysis
Smoothing & Interpolation:
Smoothing splines with variable knot amount
Polynomial interpolation
1st and 2nd order derivatives
Dew Point & RH:
Search for significant points in data
Absolute value depends on RH & temperature
Capacitive:
  T(max(ddy))
  T(dy=0)
Proximity:
  T(min(ddy))
  T(min(dy))
Plotting & Comparison:
Notes: Empirisch, immer abhängig von Steigung oder Änderungsrate der Steigung.

## Slide 30: MeasurementS – IDE CAPACITOR (4 cycles)

## Slide 31: MeasurementS – IDE CAPACITOR (2 cycles)

## Slide 32: MeasurementS – IDE CAPACITOR (2 cycles)

## Slide 33: RESULTS
Based on the average closest point of interest:
  IDE capacitor: T(max(ddy))
  Proximity sensors: T(min(dy))
Error | T_dew [°C] |  |  | RH [%] |  | 
 | IDE | VCNL4040 | VCNL36825T | IDE | VCNL4040 | VCNL36825T
Min | -0.55 | -0.99 | -0.33 | -2.97 | -5.17 | -1.48
Max | 1.18 | 0.16 | 2.9 | 4.44 | 1.22 | 8.69
Avg | 0.25 | -0.32 | 0.33 | 0.88 | -1.32 | 1.11
Std | 0.51 | 0.34 | 0.81 | 2.12 | 1.86 | 2.60